## File_description
Sheet Protocol
QC protocol
All set-ups with PMMA and RandoAlderson (RA) phantoms include the table and mattress
Sheets Canon, Siemens, GE and Philips
One sheet for QC results of each system
Doses are reported in mGy
Ratios are reported in %
Modality MSD refers to the MSD calculations performed directly on the modality such as Canon DTS, GE DoseMap and Siemens XXX
Mode is acquisition

## Protocol
Configuration | Purpose | Set-up | Phantom thickness | Projections | kV | Short Name
1 | Dosimeter/CK calibration | Free in air | NA | PA | 80 (low kV) | Air_low kV
2 | Dosimeter/CK calibration | Free in air | NA | PA | 120 (high kV) | Air_high Kv
3 | Effect of the mattress and table attenuation | On mattress | NA | PA | 80 (low kV) | Table_low kV
4 | Effect of the mattress and table attenuation | On mattress | NA | PA | 120 (high kV) | Table_high kV | Specific table data used in present document "Summary updated.xlsx" | [10 empty] | Tube voltage | Configuration | Displayed Ka,r accuracy (displayed/measured)
5 | Effect of the phantom scatter* | PMMA | 20 | PA | Automated selection | 20cm PMMA |  |  |  | Table/mattress attenuation |  |  |  | Ka,r correction factor | [5 empty] | GE | Philips | Siemens | Canon
6 | Effect of the phantom scatter *; thicker patient | PMMA | 25 | PA | Automated selection | 25cm PMMA | Name | Version | User settings | GE | Philips | Siemens | Canon | GE | Philips | Siemens | Canon | low (~80 kVp) | Free in air | 111% | 115% | 100% | 113%
7 | Effect of the phantom scatter *; thicker patient | PMMA | 30 | PA | Automated selection | 30cm PMMA | DOSE | 19.11 | Mattress thickness =4 cm |  |  |  |  | 0.92166 | 0.9009 | 0.90498 | 0.97561 | high (~120 kVp) | Free in air | 106% | 107% | 121% | 92%
8 | Effect of the phantom scatter *; thinner patient | PMMA | 15 | PA | Automated selection | 15cm PMMA | EM.dose | Modified2 | Table/mattress attenuation= 20% Table and mattress thickness =2.5 cm | 30 | 10 | 30 | 20 | 0.92166 | 0.9009 | 0.90498 | 0.97561 | Average |  | 108.5% | 111% | 110.5% | 102.5%
9 | Effect of the phantom scatter *; paediatric patient | PMMA | 10 | PA | Automated selection | 10cm PMMA | OpenSkin | Modified2 | Mattress thickness = 0 cm Table thickness = 2 cm |  |  |  |  | 0.92166 | 0.9009 | 0.90498 | 0.97561
10 | Effect of the phantom scatter * | RA | NA | PA | Automated selection | PA | RDM | 1.4.4 | Mattress thickness = 2 cm Table thickness = 0.5 cm | 30 |  | 30 |  | 0.92166 | 0.9009 | 0.90498 | 0.97561
11 | Effect of field overlap | RA | NA | PA + LAO 20 | Automated selection | PA LAO | SDM ICRP | GE | Mattress thickness = 4 cm Table equivalent thickness = 0.14 mmAl |  |  |  |  | 0.92166 | 0.9009 | 0.90498 | 0.97561
12 | Effect of field overlap | RA | NA | PA + LAO 20 + PA CRAN 15 | Automated selection | PA LAO CRAN | SDM flat | GE | Mattress thickness = 4 cm Table equivalent thickness = 0.14 mmAl |  |  |  |  | 0.92166 | 0.9009 | 0.90498 | 0.97561
13 | Effect of lateral irradiations | RA | NA | LAO 90 | Automated selection | LAT | SkinCare |  | Table/mattress attenuation= 20% | 30 | 10 | 30 | 20 | 0.92166 | 0.9009 | 0.90498 | 0.97561
100 | Repetition of high-dose clinical proc | RA | NA | See acquisition list | Automated selection | SP1 | CAATSDOSE |  | Table/mattress attenuation= 20% | 30 | 10 | 30 | 20 | 0.92166 | 0.9009 | 0.90498 | 0.97561
101 | Repetition of high-dose clinical proc | RA | NA | See acquisition list | Automated selection | SP2
102 | Repetition of high-dose clinical proc | RA | NA | See acquisition list | Automated selection | SP3
103 | Repetition of high-dose clinical proc | RA | NA | See acquisition list | Automated selection | SP4

## GE
| Measurement set-up | [11 empty] | Dose to air | Dose to skin | MSD_calculations/Ref_dose | [9 empty] | MSD calculations/TLD stat
XP exposure | Configuration | Proj | air/ phantom /table | Phant thick (cm) | Short name | Patient height (cm) | Patient weight (kg) | kV disp | kV meas | Added Filtr. | HVL | CK | QC dosimeter |  | DoseMap | CAATSDOSE | SDM flat | DOSE | RDM | EM.dose | OpenSkin | SkinCare | SDM ICRP | This calculator | Average | STDEV | Median | Min | Max | Transmission factor | air-to-skin factor
2 | 1 | PA | Free in air | NA | Air_low kV | NA | NA | 78 | 78 | 0 | 3.4 | 311 | 280 | [16 empty] | 64.45% | 1.06
3'' | 2 | PA | Free in air | NA | Air_high Kv | NA | NA | 120 | 122.5 | 0 | 5.26 | 582 | 551.3 | [16 empty] | 70.05%
4 | 3 | PA | On mattress | NA | Table_low kV | NA | NA | 80 | 81.3 | 0 | 4.01 | 395 | 229.2
5 | 4 | PA | On mattress | NA | Table_high kV | NA | NA | 120 | 124.4 | 0 | 5.71 | 749 | 497
17 | 9 | PA | PMMA | 10 | 10cm PMMA 67kVp 5.3mmAl | 140 | 45 | 67 | 67.7 | 0.3 | 5.32 | 122 | 90.4858 | 115.2326 | 138.8% | 107.9% | 91.88% | 79.98% | 95.97% | 95.88% | 120.6% | 90.28% | 90.92% | 96.33% | 101.4% | 18.22% | 95.88% | 79.98% | 138.8%
16 | 8 | PA | PMMA | 15 | 15cm PMMA 67kVp 4.7mmAl | 160 | 55 | 67 | 67 | 0.2 | 4.74 | 244 | 170.236 | 206.8038 | 154.7% | 119.9% | 102.1% | 84.68% | 106.6% | 106.8% | 134% | 100.2% | 101% | 100.1% | 112.2% | 21.01% | 106.6% | 84.68% | 154.7%
15 | 5 | PA | PMMA | 20 | 20cm PMMA 67kVp 4.8mmAl | 170 | 70 | 67 | 67 | 0.2 | 4.75 | 246 | 171.6564 | 206.8038 | 154.7% | 120.7% | 102.8% | 89.13% | 107.3% | 107.6% | 134.9% | 101% | 101.7% | 100.6% | 113.3% | 20.27% | 107.3% | 89.13% | 154.7%
14 | 6 | PA | PMMA | 25 | 25cm PMMA 71kVp 3.4mmAl | 170 | 80 | 71 | 71 | 0 | 3.35 | 461 | 293.196 | 366.6661 | 109.1% | 128.9% | 111.2% | 95.52% | 113.5% | 114.1% | 146.2% | 108.7% | 110% | 106.4% | 115.2% | 14.43% | 111.2% | 95.52% | 146.2%
13 | 7 | PA | PMMA | 30 | 30cm PMMA 93kVp 4.4mmAl | 170 | 90 | 93 | 93.7 | 0 | 4.42 | 556 | 405.132 | 495.4608 | 96.88% | 123.7% | 114.5% | 100.5% | 105.2% | 107.1% | 135.3% | 107.1% | 113.1% | 106% | 111.5% | 11.94% | 107.1% | 96.88% | 135.3%
6 | 10 | PA | RA | NA | PA | 175 | 70 | 68 | 0 | 0.2 | 0 | 320 | NA | 366.4427 | 114.6% | 90.81% | 101.9% | 65.39% | 89.77% | 84.95% | 101.2% | 75.71% | 96.25% | 78.05% | 91.18% | 14.73% | 90.81% | 65.39% | 114.6%
7 | 11 | PA + LAO 20 | RA | NA | PA LAO | 175 | 70 | 67 | 0 | 0.2 | 0 | 301 | NA | 360.9186 | 116.4% | 90.88% | 99.77% | 61.29% | 95.24% | 83.57% | 98.14% | 75.97% | 95.64% | 82.57% | 90.76% | 15.71% | 95.24% | 61.29% | 116.4%
8 | 12 | PA + LAO 20 + PA CRAN 15 | RA | NA | PA LAO CRAN | 175 | 70 | 67 | 0 | 0.2 | 0 | 296 | NA | 316.3807 | 126.4% | 100.8% | 110.3% | 69.92% | 104.3% | 92.78% | 107.5% | 83.35% | 102.2% | 90.4% | 99.74% | 16.27% | 102.2% | 69.92% | 126.4%
9 | 13 | LAO 90 | RA | NA | LAT | 175 | 70 | 71 | 0 | 0.2 | 0 | 311 | NA | 480.7828 | 120.6% | 75.48% | NA | 101.6% | 67.48% | NA | NA | 113.4% | 207.4% | 82.37% | 114.3% | 50.13% | 107.5% | 67.48% | 207.4%
| Measurement set-up | [11 empty] | Dose to air | Dose to skin | MSD calculations/Film | [9 empty] | MSD calculations/Film stat
10 | 100 | See acquisition list | RA | NA | SP1 | 170 | 75 | NA | NA | NA | 0 | 1000 | NA | 662.5816 | 114.1% | 78.07% | 128.6% | 70.94% | 98.4% | 117.2% | 91.97% | 99.68% | 127.1% | 99.31% | 102.9% | 20.5% | 99.68% | 70.94% | 128.6%
11 | 101 | See acquisition list | RA | NA | SP2 | 170 | 75 | NA | NA | NA | 0 | 999 | NA | 744.4295 | 101.5% | 80.22% | 111.4% | 66.86% | 95.66% | 102.8% | 103.6% | 85.91% | 105.4% | 92.02% | 94.82% | 14.35% | 101.5% | 66.86% | 111.4%
12 | 102 | See acquisition list | RA | NA | SP3 | 170 | 75 | NA | NA | NA | 0 | 999 | NA | 731.5291 | 90.71% | 80.41% | 105.6% | 59.22% | 90.02% | 90.84% | 94.22% | 79.59% | 100.5% | 90.36% | 87.91% | 13.61% | 90.71% | 59.22% | 105.6%

## Siemens
| Measurement set-up | [11 empty] | Dose to air | Dose to skin | MSD_calculations/Ref_dose | [9 empty] | MSD calculations/TLD stat
XP exposure | Configuration | Proj | air/ phantom /table | Phant thick (cm) | Short name | Patient height (cm) | Patient weight (kg) | kV | Total filtration mmAl | Added Filtr. | HVL | CK | QC dosimeter |  | CareMonitor | CAATSDOSE | SDM flat | DOSE | RDM | EM.dose | OpenSkin | SkinCare | SDM ICRP | this calculator | Average | Stdev | Median | Min | Max | Transmission factor | air-to-skin factor
2 | 1 | PA | Free in air | NA | Air_low kV | NA | NA | 81 | 0 | 0 | 2.85 | 313 | 311.9 | [16 empty] | 53.85% | 1.06
3 | 2 | PA | Free in air | NA | Air_high Kv | NA | NA | 109 | 0 | 0 | 3.86 | 1392 | 1146 | [16 empty] | 69.46%
4 | 3 | PA | On mattress | NA | Table_low kV | NA | NA | 81 | 0 | 0 | 3.26 | 614 | 329.5
7 | 4 | PA | On mattress | NA | Table_high kV | NA | NA | 123.4 | 0 | 0 | 4.79 | 1562 | 893.2
13 | 9 | PA | PMMA | 10 | 10cm PMMA 92kVp 3.8mmAl | 140 | 45 | 92 | 0 | 0 | 3.8 | 512 | 318 | 354.7597 | 176.2% | 175.2% | 151.3% | 119.9% | 144.5% | 158.7% | 200.8% | 126.8% | 147.7% | 131.9% | 155.7% | 25.41% | 151.3% | 119.9% | 200.8%
12 | 8 | PA | PMMA | 15 | 15cm PMMA 100kVp 4.2mmAl | 160 | 55 | 100 | 0 | 0 | 4.15 | 637 | 433.434 | 475.2363 | 158.4% | 164.3% | 146.7% | 118.1% | 137.5% | 147.5% | 186.4% | 119.8% | 143.3% | 118% | 146.9% | 21.4% | 146.7% | 118.1% | 186.4%
11 | 5 | PA | PMMA | 20 | 20cm PMMA 108kVp 4.3mmAl | 170 | 70 | 108 | 0 | 0 | 4.32 | 681 | 514.1 | 584.3726 | 142% | 144% | 131.6% | 106.9% | 119.6% | 128.3% | 162.1% | 104.9% | 128.5% | 103.7% | 129.8% | 18.22% | 128.5% | 104.9% | 162.1%
10 | 6 | PA | PMMA | 25 | 25cm PMMA 109kVp 4.2mmAl | 170 | 80 | 109 | 0 | 0 | 4.2 | 1442 | 1134.2 | 1338.0146 | 138.6% | 133.1% | 121.6% | 99.42% | 110.6% | 118.7% | 149.9% | 97.06% | 118.8% | 95.66% | 120.9% | 17.51% | 118.8% | 97.06% | 149.9%
8 | 7 | PA | PMMA | 30 | 30cm PMMA 125kVp 4.8mmAl | 170 | 90 | 125 | 0 | 0 | 4.8 | 1729.31 | 1403.44 | 1743.6951 | 134.8% | 123.5% | 117.4% | 97.57% | 101.8% | 109.4% | 138.1% | 90.28% | 115.5% | 78% | 114.3% | 16.28% | 115.5% | 90.28% | 138.1%
16 | 10 | PA | RA | NA | PA | 175 | 70 | 77.3 | 0 | 0 | 0 | 305 | NA | 337.023 | 90.2% | 99% | 91.14% | 77.87% | 85.64% | 83.88% | 106.9% | 72.84% | 90.15% | 122.2% | 88.62% | 10.3% | 90.15% | 72.84% | 106.9%
17 | 11 | PA + LAO 20 | RA | NA | PA LAO | 175 | 70 | 77.3 | 0 | 0 | 0 | 425 | NA | 527.2657 | 81.93% | 90.21% | 81.52% | 68.65% | 76.33% | 77.49% | 95.43% | 66.23% | 79.8% | 92.22% | 79.73% | 9.259% | 79.8% | 66.23% | 95.43%
19 | 12 | PA + LAO 20 + PA CRAN 15 | RA | NA | PA LAO CRAN | 175 | 70 | 77.3 | 0 | 0 | 0 | 655 | NA | 760.436 | 87.19% | 96.63% | 86.94% | 73.78% | 81.33% | 83.1% | 101.5% | 70.5% | 83.7% | 98.1% | 84.96% | 9.82% | 83.7% | 70.5% | 101.5%
20 | 13 | LAO 90 | RA | NA | LAT | 175 | 70 | 93 | 0 | 0 | 0 | 495 | NA | 890.9815 | 81.6% | 78.34% | NA | 88.37% | 33.66% | NA | NA | 79.19% | 122% | 80.14% | 80.52% | 28.23% | 80.39% | 33.66% | 122%
21 | 100 | See acquisition list | RA | NA | SP1 | 170 | 75 | 0 | 0 | 0 | 0 | 1006 | NA | 874.5 | 72.38% | 69.9% | 60.69% | 43.46% | 61.47% | 65.83% | 78.59% | 77.6% | 71.11% | 72.04% | 66.78% | 10.76% | 69.9% | 43.46% | 78.59%
22 | 101 | See acquisition list | RA | NA | SP2 | 170 | 75 | 0 | 0 | 0 | 0 | 999 | NA | 907.36 | 92.47% | 69.21% | 74.86% | 51.86% | 65.9% | 64.49% | 77.58% | 75.4% | 72.6% | 87.4% | 71.6% | 11.04% | 72.6% | 51.86% | 92.47%
23 | 102 | See acquisition list | RA | NA | SP3 | 170 | 75 | 0 | 0 | 0 | 0 | 1002 | NA | 846.94 | 80.05% | 48.24% | 71.63% | 56.63% | 67.88% | 73.96% | 82.84% | 65.35% | 68.58% | 89.14% | 68.35% | 10.84% | 68.58% | 48.24% | 82.84%

## Canon
| Measurement set-up | [11 empty] | Dose to air | Dose to skin | MSD_calculations/Ref_dose | [8 empty] | MSD calculations/TLD stat
XP exposure | Configuration | Proj | air/ phantom /table | Phant thick (cm) | Short name | Patient height (cm) | Patient weight (kg) | kV | Total filtration mmAl | Added Filtr. | HVL | CK | QC dosimeter |  | DTS | CAATSDOSE | SDM | DOSE | RDM | EM.dose | OpenSkin | SkinCare | This calculator | Average | SDEV | Median | Min | Max | Transmission factor | air-to-skin factor
1 | 1 | PA | Free in air | NA | Air_low kV | NA | NA | 80 | 0 | 0 | 5.25 | 494.07 | 436.95 | [15 empty] | 84.85% | 1.06
2 | 2 | PA | Free in air | NA | Air_high Kv | NA | NA | 120 | 0 | 0 | 7.2 | 497.04 | 540.7 | [15 empty] | 78.29%
3 | 3 | PA | On mattress | NA | Table_low kV | NA | NA | 80 | 0 | 0 | 5.25 | 463.9 | 348.13
4 | 4 | PA | On mattress | NA | Table_high kV | NA | NA | 120 | 0 | 0 | 7.2 | 520.5 | 443.28
10 | 9 | PA | PMMA | 10 | 10cm PMMA 74kVp 4.6mmAl | 140 | 40 | 74 | 0 | 0 | 4.6 | 118.19 | 90.895 | 140.3692 | 94.75% | 104.8% | NA |  | NA | 91.05% | 102.9% | 94.52% | 88.34% | 97.6% | 5.922% | 94.75% | 91.05% | 104.8%
8 | 8 | PA | PMMA | 15 | 15cm PMMA 74kVp 4.7mmAl | 150 | 55 | 74 | 0 | 0 | 4.7 | 193.34 | 151.686 | 235.7809 | 86.1% | 102% | NA |  | NA | 84.82% | 100.3% | 92.27% | 85.67% | 93.09% | 7.889% | 92.27% | 84.82% | 102%
7 | 5 | PA | PMMA | 20 | 20cm PMMA 83kVp 7.2mmAl | 170 | 66 | 83 | 0 | 0 | 7.15 | 430.9 | 359.7852 | 564.5458 | 81.84% | 96.19% | NA |  | NA | 82.43% | 93.35% | 88.31% | 80.77% | 88.42% | 6.402% | 88.31% | 81.84% | 96.19%
5' | 6 | PA | PMMA | 25 | 25cm PMMA 105kVp 6.3mmAl | 170 | 66 | 105 | 7.6 | 0 | 6.34 | 264.31 | 250.5204 | 371.8376 | 141.5% | 99.95% | NA |  | NA | 84.75% | 99.06% | 92.62% | 98.7% | 103.6% | 22.04% | 99.06% | 84.75% | 141.5%
6 | 7 | PA | PMMA | 30 | 30cm PMMA 125kVp 7.2mmAl | 170 | 70 | 125 | 0 | 0 | 7.15 | 382.12 | 356.478 | 543.5263 | 155.1% | 89.51% | NA |  | NA | 75.03% | 88.05% | 83.11% | 95.49% | 98.16% | 32.33% | 88.05% | 75.03% | 155.1%
11 | 10 | PA | RA | NA | PA | 175 | 70 | 75 | 0 | 0 | 0 | 519.77 | NA | 442.2623 | 111.5% | 128.3% | NA |  | NA | 109.4% | 123.3% | 118.2% | 114.9% | 118.1% | 7.892% | 118.2% | 109.4% | 128.3%
15 | 11 | PA + LAO 20 | RA | NA | PA LAO | 190 | 85 | 74 | 0 | 0 | 0 | 1052.91 | NA | 1343.9509 | 105.9% | 84.6% | NA |  | NA | 72.08% | 79.53% | 77.02% | 81.1% | 83.82% | 13.13% | 79.53% | 72.08% | 105.9%
16 | 12 | PA + LAO 20 + PA CRAN 15 | RA | NA | PA LAO CRAN | 190 | 85 | 0 | 0 | 0 | 0 | 1315.85 | NA | 1924.0004 | 133.6% | 71.64% | NA |  | NA | 61.31% | 67.65% | 65.51% | 71.73% | 79.94% | 30.22% | 67.65% | 61.31% | 133.6%
17 | 13 | LAO 90 | RA | NA | LAT | 190 | 85 | 73 | 0 | 0 | 0 | 369.49 | NA | 704.3951 | 160.8% | 85.88% | NA |  | NA | NA | NA | 61.63% | 69.42% | 102.8% | 51.72% | 85.88% | 61.63% | 160.8%
| Measurement set-up | [11 empty] | Dose to air | Dose to skin | MSD calculations/Film | [8 empty] | MSD calculations/Film stat
20 | 100 | See acquisition list | RA | NA | SP1 | 175 | 75 | 0 | 0 | 0 | 0 | 999.979 | NA | 1039.86 | 84.53% | 81.16% | NA | 109.4% | NA | 72.99% | 82.76% | 74.12% | 71.45% | 84.16% | 13.21% | 81.96% | 72.99% | 109.4%
21 | 101 | See acquisition list | RA | NA | SP2 | 175 | 75 | 0 | 0 | 0 | 0 | 1001.454 | NA | 1066.36 | 84.31% | 83.68% | NA | 95.66% | NA | 74.02% | 78.95% | 79.14% | 74.65% | 82.62% | 7.398% | 81.41% | 74.02% | 95.66%
22 | 102 | See acquisition list | RA | NA | SP3 | 175 | 75 | 0 | 0 | 0 | 0 | 999.945 | NA | 969.9 | 71.55% | 72.29% | NA | 83.75% | NA | 83.19% | 86.16% | 88.12% | 77.84% | 80.84% | 7.134% | 83.47% | 71.55% | 88.12%

## Philips
| Measurement set-up | [11 empty] | Dose to air | Dose to skin | MSD_calculations/Ref_dose | [9 empty] | MSD calculations/TLD stat
XP exposure | Configuration | Proj | air/ phantom /table | Phant thick (cm) | Short name | Patient height (cm) | Patient weight (kg) | kV | Total filtration mmAl | Added Filtr. | HVL | CK | QC dosimeter |  | Modality | CAATSDOSE | SDM flat | DOSE | RDM | EM.dose | OpenSkin | SkinCare | SDM ICRP | This calculator | Average | Stdev | Median | Min | Max | Transmission factor | air-to-skin factor
1 | 1 | PA | Free in air | NA | Air_low kV | NA | NA | 68 | 68.2 | 0 | 4.32 | 114 | 98.93 | [16 empty] | 88.63% | 1.06
2 | 2 | PA | Free in air | NA | Air_high Kv | NA | NA | 125 | 131 | 0 | 7.61 | 213 | 199.6 | [16 empty] | 89.83%
3 | 3 | PA | On mattress | NA | Table_low kV | NA | NA | 68 | 68 | 0 | 4.51 | 87 | 66.916
4 | 4 | PA | On mattress | NA | Table_high kV | NA | NA | 125 | 131 | 0 | 7.79 | 220 | 185.2
9 | 9 | PA | PMMA | 10 | 10cm PMMA 63kVp 4.1mmAl | 140 | 45 | 63 | 63.9 | 0 | 4.07 | 101 | 82.3302 | 108.9082 | NA | 142.1% | 103.2% | 140.6% | NA | 114.5% | 130.1% | 99.58% | 97.04% | 84.47% | 118.2% | 19.37% | 114.5% | 97.04% | 142.1%
8 | 8 | PA | PMMA | 15 | 15cm PMMA 67kVp 4.4mmAl | 160 | 55 | 67 | 67.6 | 0 | 4.37 | 155 | 133.984 | 177.1783 | NA | 135.4% | 99.4% | 137.3% | NA | 108.7% | 123.1% | 95.53% | 93.48% | 87.48% | 113.3% | 18.63% | 108.7% | 93.48% | 137.3%
5'' | 5 | PA | PMMA | 20 | 20cm PMMA 72kVp 4.7mmAl | 170 | 70 | 72 | 72.4 | 0 | 4.724 | 161 | 142.3686 | 185.6864 | NA | 135.6% | 101% | 135.8% | NA | 108.1% | 122.1% | 96.06% | 95.02% | 86.71% | 113.4% | 17.75% | 108.1% | 95.02% | 135.8%
6 | 6 | PA | PMMA | 25 | 25cm PMMA 78kVp 5.2mmAl | 170 | 80 | 78 | 81.6 | 0 | 5.22 | 212 | 186.454 | 240.037 | NA | 133.8% | 100.2% | 135.1% | NA | 104.7% | 116.8% | 96.27% | 94.42% | 84.99% | 111.6% | 17.22% | 104.7% | 94.42% | 135.1%
7 | 7 | PA | PMMA | 30 | 30cm PMMA 96kVp 6.1mmAl | 170 | 90 | 96 | 98.3 | 0 | 6.12 | 278 | 247.192 | 312.391 | NA | 129.7% | 100.7% | 132.7% | NA | 100.9% | 114.6% | 95.71% | 95.1% | 93.47% | 109.9% | 15.9% | 100.9% | 95.1% | 132.7%
10' | 10 | PA | RA | NA | PA | 170 | 75 | NA | NA | 0 | NA | 107 | NA | 134.1837 | NA | 118.4% | 85.17% | 120.9% | NA | 92.9% | 104.9% | 83.84% | 84.88% | 70.8% | 98.7% | 16.05% | 92.9% | 83.84% | 120.9%
11 | 11 | PA + LAO 20 | RA | NA | PA LAO | 170 | 75 | NA | NA | 0 | NA | 104 | NA | 110.6283 | NA | 143.2% | 97.02% | 138.4% | NA | 113.6% | 71.45% | 99.86% | 96.11% | 84.07% | 108.5% | 25.37% | 99.86% | 71.45% | 143.2%
12 | 12 | PA + LAO 20 + PA CRAN 15 | RA | NA | PA LAO CRAN | 170 | 75 | NA | NA | 0 | NA | 150 | NA | 158.3164 | NA | 144.2% | 99.05% | 136.6% | NA | 114.6% | 83.16% | 100.5% | 97.03% | 84.01% | 110.7% | 22.33% | 100.5% | 83.16% | 144.2%
13 | 13 | LAO 90 | RA | NA | LAT | 170 | 75 | NA | NA | 0 | NA | 125 | NA | 189.7003 | NA | 119.4% | NA | 109.2% | NA | NA | NA | 122.9% | 150% | 83.82% | 125.4% | 17.39% | 121.2% | 109.2% | 150%
| Measurement set-up | [11 empty] | Dose to air | Dose to skin | MSD calculations/Film | [9 empty] | MSD calculations/Film stat
14 | 100 | See acquisition list | RA | NA | SP1 | 170 | 75 | NA | NA | NA | NA | 500 | NA | 306.0328 | NA | 155% | 112.8% | 144.2% | NA | 126.8% | 140.3% | 145.1% | 110.5% | 104.9% | 133.5% | 17.13% | 140.3% | 110.5% | 155%
15 | 101 | See acquisition list | RA | NA | SP2 | 170 | 75 | NA | NA | NA | NA | 500 | NA | 307.9273 | NA | 115.9% | 114.6% | 114.1% | NA | 149.1% | 103.4% | 128% | 113.1% | 103.6% | 119.7% | 14.81% | 114.6% | 103.4% | 149.1%
16 | 102 | See acquisition list | RA | NA | SP3 | 170 | 75 | NA | NA | NA | NA | 501 | NA | 274.2965 | NA | 85.81% | 121% | 139.1% | NA | 137.7% | 110.8% | 123.8% | 117.2% | 105.7% | 119.3% | 18.03% | 121% | 85.81% | 139.1%